## Notes
Tab colours
populated sheets used to determine consensus δ(13C) values and uncertainty for BEET-1, GALT-1, and FRUT-1
reported measurements of BEET-1, GALT-1, FRUT-1, IAEA-CH-6, USGS40, and USGS62 from each participating laboratory
measurements performed at NRC for homogeneity, thermal stability and temporal stability
effect of the oxygen-17 correction on NRC measurement results

## BEET-1
Consensus value
Calculations shown here were performed using spreadsheet published by | [9 empty] | mean | u | U95%
Meija and Chartrand (2017) Anal Bioanal Chem, doi: 10.1007/s00216-017-0659-1 | [9 empty] | -26.019 | 0.045 | 0.087
(1) Data for secondary reference materials |  |  |  | (2) Measurements of the secondary reference materials |  |  |  |  | (3) Measurements of the sample
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13
Calibrator | x | u(x) | df | Lab | Calibrator | y | s(y) | N | Lab | y | s(y) | N | EXPLANATION OF COLUMNS
IAEA-CH-6 | -10.449 | 0.033 | 100 | A | IAEA-CH-6 | -10.915 | 0.093 | 4 | A | -26.606 | 0.106 | 12 | 1 | Name of the calibrators
USGS40 | -26.39 | 0.04 | 100 | B | IAEA-CH-6 | -10.575 | 0.085 | 4 | B | -26.066 | 0.091 | 12 | 2 | True isotope delta value assigned to the calibrator
USGS62 | -14.79 | 0.07 | 100 | C | IAEA-CH-6 | -10.553 | 0.074 | 4 | C | -26.056 | 0.21 | 12 | 3 | Standard uncertainty assigned to the value in 1
|  |  |  | D | IAEA-CH-6 | -10.45 | 0.014 | 4 | D | -26.054 | 0.059 | 12 | 4 | Degrees-of-freedom associated with the standard uncertainty in 2
|  |  |  | E | IAEA-CH-6 | -10.455 | 0.066 | 4 | E | -26.037 | 0.037 | 12 | 5 | Laboratory name
|  |  |  | F | IAEA-CH-6 | -10.465 | 0.071 | 4 | F | -25.987 | 0.073 | 12 | 6 | Calibrator name
|  |  |  | G | IAEA-CH-6 | -10.418 | 0.079 | 4 | G | -26.031 | 0.084 | 12 | 7 | Measured average isotope delta value for calibrator
|  |  |  | H | IAEA-CH-6 | -10.445 | 0.056 | 4 | H | -26.042 | 0.059 | 12 | 8 | Standard deviation of N replicate measurements whose mean is reported in 7
|  |  |  | I | IAEA-CH-6 | -10.44 | 0.045 | 4 | I | -26.035 | 0.08 | 12 | 9 | Number of replicate measurements whose average is reported in 7
|  |  |  | A | USGS40 | -27.095 | 0.064 | 4 |  |  |  |  | 10 | Laboratory name
|  |  |  | B | USGS40 | -26.373 | 0.071 | 4 |  |  |  |  | 11 | Measured average isotope delta values for sample
|  |  |  | C | USGS40 | -26.32 | 0.268 | 4 |  |  |  |  | 12 | Standard deviation of N replicate measurements whose mean is reported in 11
|  |  |  | D | USGS40 | -26.43 | 0.014 | 4 |  |  |  |  | 13 | Number of replicate measurements whose average is reported in 11
|  |  |  | E | USGS40 | -26.403 | 0.061 | 4
|  |  |  | F | USGS40 | -26.348 | 0.083 | 4 |  |  |  |  | 10000 | <- Number of Monte Carlo simulations
|  |  |  | G | USGS40 | -26.388 | 0.075 | 4
|  |  |  | H | USGS40 | -26.363 | 0.049 | 4
|  |  |  | I | USGS40 | -26.403 | 0.025 | 4
|  |  |  | A | USGS62 | -15.295 | 0.029 | 4
|  |  |  | B | USGS62 | -14.925 | 0.013 | 4
|  |  |  | C | USGS62 | -14.99 | 0.095 | 4
|  |  |  | D | USGS62 | -14.798 | 0.013 | 4
|  |  |  | E | USGS62 | -14.846 | 0.03 | 5
|  |  |  | F | USGS62 | -14.675 | 0.182 | 4
|  |  |  | G | USGS62 | -14.84 | 0.092 | 4
|  |  |  | H | USGS62 | -14.8 | 0.024 | 4
|  |  |  | I | USGS62 | -14.813 | 0.028 | 4

## GALT-1
Consensus value
Calculations shown here were performed using spreadsheet published by | [9 empty] | mean | u | U95%
Meija and Chartrand (2017) Anal Bioanal Chem, doi: 10.1007/s00216-017-0659-1 | [9 empty] | -21.410 | 0.032 | 0.062
(1) Data for secondary reference materials |  |  |  | (2) Measurements of the secondary reference materials |  |  |  |  | (3) Measurements of the sample
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13
Calibrator | x | u(x) | df | Lab | Calibrator | y | s(y) | N | Lab | y | s(y) | N | EXPLANATION OF COLUMNS
IAEA-CH-6 | -10.449 | 0.033 | 100 | A | IAEA-CH-6 | -10.915 | 0.093 | 4 | A | -21.951 | 0.061 | 12 | 1 | Name of the calibrators
USGS40 | -26.39 | 0.04 | 100 | B | IAEA-CH-6 | -10.575 | 0.085 | 4 | B | -21.481 | 0.09 | 12 | 2 | True isotope delta value assigned to the calibrator
USGS62 | -14.79 | 0.07 | 100 | C | IAEA-CH-6 | -10.553 | 0.074 | 4 | C | -21.371 | 0.184 | 12 | 3 | Standard uncertainty assigned to the value in 1
|  |  |  | D | IAEA-CH-6 | -10.45 | 0.014 | 4 | D | -21.473 | 0.022 | 12 | 4 | Degrees-of-freedom associated with the standard uncertainty in 2
|  |  |  | E | IAEA-CH-6 | -10.455 | 0.066 | 4 | E | -21.443 | 0.033 | 13 | 5 | Laboratory name
|  |  |  | F | IAEA-CH-6 | -10.465 | 0.071 | 4 | F | -21.397 | 0.106 | 12 | 6 | Calibrator name
|  |  |  | G | IAEA-CH-6 | -10.418 | 0.079 | 4 | G | -21.389 | 0.05 | 12 | 7 | Measured average isotope delta value for calibrator
|  |  |  | H | IAEA-CH-6 | -10.445 | 0.056 | 4 | H | -21.389 | 0.046 | 12 | 8 | Standard deviation of N replicate measurements whose mean is reported in 7
|  |  |  | I | IAEA-CH-6 | -10.44 | 0.045 | 4 | I | -21.415 | 0.025 | 12 | 9 | Number of replicate measurements whose average is reported in 7
|  |  |  | A | USGS40 | -27.095 | 0.064 | 4 |  |  |  |  | 10 | Laboratory name
|  |  |  | B | USGS40 | -26.373 | 0.071 | 4 |  |  |  |  | 11 | Measured average isotope delta values for sample
|  |  |  | C | USGS40 | -26.32 | 0.268 | 4 |  |  |  |  | 12 | Standard deviation of N replicate measurements whose mean is reported in 11
|  |  |  | D | USGS40 | -26.43 | 0.014 | 4 |  |  |  |  | 13 | Number of replicate measurements whose average is reported in 11
|  |  |  | E | USGS40 | -26.403 | 0.061 | 4
|  |  |  | F | USGS40 | -26.348 | 0.083 | 4 |  |  |  |  | 10000 | <- Number of Monte Carlo simulations
|  |  |  | G | USGS40 | -26.388 | 0.075 | 4
|  |  |  | H | USGS40 | -26.363 | 0.049 | 4
|  |  |  | I | USGS40 | -26.403 | 0.025 | 4
|  |  |  | A | USGS62 | -15.295 | 0.029 | 4
|  |  |  | B | USGS62 | -14.925 | 0.013 | 4
|  |  |  | C | USGS62 | -14.99 | 0.095 | 4
|  |  |  | D | USGS62 | -14.798 | 0.013 | 4
|  |  |  | E | USGS62 | -14.846 | 0.03 | 5
|  |  |  | F | USGS62 | -14.675 | 0.182 | 4
|  |  |  | G | USGS62 | -14.84 | 0.092 | 4
|  |  |  | H | USGS62 | -14.8 | 0.024 | 4
|  |  |  | I | USGS62 | -14.813 | 0.028 | 4

## FRUT-1
Consensus value
Calculations shown here were performed using spreadsheet published by | [9 empty] | mean | u | U95%
Meija and Chartrand (2017) Anal Bioanal Chem, doi: 10.1007/s00216-017-0659-1 | [9 empty] | -10.984 | 0.036 | 0.070
(1) Data for secondary reference materials |  |  |  | (2) Measurements of the secondary reference materials |  |  |  |  | (3) Measurements of the sample
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13
Calibrator | x | u(x) | df | Lab | Calibrator | y | s(y) | N | Lab | y | s(y) | N | EXPLANATION OF COLUMNS
IAEA-CH-6 | -10.449 | 0.033 | 100 | A | IAEA-CH-6 | -10.915 | 0.093 | 4 | A | -11.498 | 0.065 | 12 | 1 | Name of the calibrators
USGS40 | -26.39 | 0.04 | 100 | B | IAEA-CH-6 | -10.575 | 0.085 | 4 | B | -11.128 | 0.068 | 11 | 2 | True isotope delta value assigned to the calibrator
USGS62 | -14.79 | 0.07 | 100 | C | IAEA-CH-6 | -10.553 | 0.074 | 4 | C | -11.12 | 0.106 | 12 | 3 | Standard uncertainty assigned to the value in 1
|  |  |  | D | IAEA-CH-6 | -10.45 | 0.014 | 4 | D | -10.973 | 0.076 | 12 | 4 | Degrees-of-freedom associated with the standard uncertainty in 2
|  |  |  | E | IAEA-CH-6 | -10.455 | 0.066 | 4 | E | -11.001 | 0.042 | 11 | 5 | Laboratory name
|  |  |  | F | IAEA-CH-6 | -10.465 | 0.071 | 4 | F | -11.032 | 0.18 | 12 | 6 | Calibrator name
|  |  |  | G | IAEA-CH-6 | -10.418 | 0.079 | 4 | G | -11.009 | 0.034 | 12 | 7 | Measured average isotope delta value for calibrator
|  |  |  | H | IAEA-CH-6 | -10.445 | 0.056 | 4 | H | -10.998 | 0.057 | 12 | 8 | Standard deviation of N replicate measurements whose mean is reported in 7
|  |  |  | I | IAEA-CH-6 | -10.44 | 0.045 | 4 | I | -10.953 | 0.033 | 12 | 9 | Number of replicate measurements whose average is reported in 7
|  |  |  | A | USGS40 | -27.095 | 0.064 | 4 |  |  |  |  | 10 | Laboratory name
|  |  |  | B | USGS40 | -26.373 | 0.071 | 4 |  |  |  |  | 11 | Measured average isotope delta values for sample
|  |  |  | C | USGS40 | -26.32 | 0.268 | 4 |  |  |  |  | 12 | Standard deviation of N replicate measurements whose mean is reported in 11
|  |  |  | D | USGS40 | -26.43 | 0.014 | 4 |  |  |  |  | 13 | Number of replicate measurements whose average is reported in 11
|  |  |  | E | USGS40 | -26.403 | 0.061 | 4
|  |  |  | F | USGS40 | -26.348 | 0.083 | 4 |  |  |  |  | 10000 | <- Number of Monte Carlo simulations
|  |  |  | G | USGS40 | -26.388 | 0.075 | 4
|  |  |  | H | USGS40 | -26.363 | 0.049 | 4
|  |  |  | I | USGS40 | -26.403 | 0.025 | 4
|  |  |  | A | USGS62 | -15.295 | 0.029 | 4
|  |  |  | B | USGS62 | -14.925 | 0.013 | 4
|  |  |  | C | USGS62 | -14.99 | 0.095 | 4
|  |  |  | D | USGS62 | -14.798 | 0.013 | 4
|  |  |  | E | USGS62 | -14.846 | 0.03 | 5
|  |  |  | F | USGS62 | -14.675 | 0.182 | 4
|  |  |  | G | USGS62 | -14.84 | 0.092 | 4
|  |  |  | H | USGS62 | -14.8 | 0.024 | 4
|  |  |  | I | USGS62 | -14.813 | 0.028 | 4

## Lab A
Sample Identity | Sample Name | Replicate Number | Reported d(13C), ‰
IAEA-CH-6 | A | 1 | -11.03
IAEA-CH-6 | A | 2 | -10.95
IAEA-CH-6 | A | 3 | -10.84
IAEA-CH-6 | A | 4 | -10.84
USGS40 | B | 1 | -27.19
USGS40 | B | 2 | -27.07
USGS40 | B | 3 | -27.07
USGS40 | B | 4 | -27.05
USGS62 | C | 1 | -15.29
USGS62 | C | 2 | -15.26
USGS62 | C | 3 | -15.3
USGS62 | C | 4 | -15.33
BEET-1 | D | 1 | -26.87
BEET-1 | D | 2 | -26.59
BEET-1 | D | 3 | -26.58
BEET-1 | D | 4 | -26.66
BEET-1 | I | 1 | -26.67
BEET-1 | I | 2 | -26.49
BEET-1 | I | 3 | -26.56
BEET-1 | I | 4 | -26.52
BEET-1 | K | 1 | -26.7
BEET-1 | K | 2 | -26.5
BEET-1 | K | 3 | -26.55
BEET-1 | K | 4 | -26.58
FRUT-1 | F | 1 | -11.36
FRUT-1 | F | 2 | -11.53
FRUT-1 | F | 3 | -11.58
FRUT-1 | F | 4 | -11.51
FRUT-1 | H | 1 | -11.4
FRUT-1 | H | 2 | -11.52
FRUT-1 | H | 3 | -11.47
FRUT-1 | H | 4 | -11.47
FRUT-1 | J | 1 | -11.55
FRUT-1 | J | 2 | -11.56
FRUT-1 | J | 3 | -11.51
FRUT-1 | J | 4 | -11.52
GALT-1 | E | 1 | -21.91
GALT-1 | E | 2 | -22
GALT-1 | E | 3 | -22
GALT-1 | E | 4 | -21.97
GALT-1 | G | 1 | -22.07
GALT-1 | G | 2 | -21.94
GALT-1 | G | 3 | -21.97
GALT-1 | G | 4 | -21.97
GALT-1 | L | 1 | -21.84
GALT-1 | L | 2 | -21.88
GALT-1 | L | 3 | -21.95
GALT-1 | L | 4 | -21.91

## Lab B
Sample Identity | Sample Name | Replicate Number | Reported d(13C), ‰
IAEA-CH-6 | A | 1 | -10.58
IAEA-CH-6 | A | 2 | -10.49
IAEA-CH-6 | A | 3 | -10.69
IAEA-CH-6 | A | 4 | -10.54
USGS40 | B | 1 | -26.31
USGS40 | B | 2 | -26.4
USGS40 | B | 3 | -26.46
USGS40 | B | 4 | -26.32
USGS62 | C | 1 | -14.94
USGS62 | C | 2 | -14.92
USGS62 | C | 3 | -14.91
USGS62 | C | 4 | -14.93
BEET-1 | D | 1 | -26.22
BEET-1 | D | 2 | -26.16
BEET-1 | D | 3 | -25.92
BEET-1 | D | 4 | -25.95
BEET-1 | I | 1 | -26.12
BEET-1 | I | 2 | -26.14
BEET-1 | I | 3 | -26.02
BEET-1 | I | 4 | -26.08
BEET-1 | K | 1 | -25.98
BEET-1 | K | 2 | -26
BEET-1 | K | 3 | -26.1
BEET-1 | K | 4 | -26.1
FRUT-1 | F | 1 | -11.07
FRUT-1 | F | 2 | -11.18
FRUT-1 | F | 3 | -11.13
FRUT-1 | F | 4 | -11.11
FRUT-1 | H | 1 | -11.1
FRUT-1 | H | 2 | -11.1
FRUT-1 | H | 3 | -11.1
FRUT-1 | H | 4 | -11.28
FRUT-1 | J | 1 | -11.02
FRUT-1 | J | 2 | -11.23
FRUT-1 | J | 3 | -11.25
FRUT-1 | J | 4 | -11.12
GALT-1 | E | 1 | -21.48
GALT-1 | E | 2 | -21.46
GALT-1 | E | 3 | -21.5
GALT-1 | E | 4 | -21.63
GALT-1 | G | 1 | -21.34
GALT-1 | G | 2 | -21.47
GALT-1 | G | 3 | -21.42
GALT-1 | G | 4 | -21.64
GALT-1 | L | 1 | -21.42
GALT-1 | L | 2 | -21.47
GALT-1 | L | 3 | -21.55
GALT-1 | L | 4 | -21.39
This laboratory reported a 10 mV N2 peak, so this measurment was discarded

## Lab C
Sample Identity | Sample Name | Replicate Number | Reported d(13C), ‰
IAEA-CH-6 | A | 1 | -10.51
IAEA-CH-6 | A | 2 | -10.47
IAEA-CH-6 | A | 3 | -10.62
IAEA-CH-6 | A | 4 | -10.61
USGS40 | B | 1 | -26.06
USGS40 | B | 2 | -26.12
USGS40 | B | 3 | -26.58
USGS40 | B | 4 | -26.52
USGS62 | C | 1 | -15
USGS62 | C | 2 | -14.92
USGS62 | C | 3 | -14.92
USGS62 | C | 4 | -15.12
BEET-1 | D | 1 | -25.69
BEET-1 | D | 2 | -25.93
BEET-1 | D | 3 | -25.92
BEET-1 | D | 4 | -25.67
BEET-1 | I | 1 | -26.11
BEET-1 | I | 2 | -26.16
BEET-1 | I | 3 | -26.18
BEET-1 | I | 4 | -26.24
BEET-1 | K | 1 | -26.27
BEET-1 | K | 2 | -26.08
BEET-1 | K | 3 | -26.13
BEET-1 | K | 4 | -26.29
FRUT-1 | F | 1 | -11.04
FRUT-1 | F | 2 | -11.06
FRUT-1 | F | 3 | -10.97
FRUT-1 | F | 4 | -11.1
FRUT-1 | H | 1 | -10.94
FRUT-1 | H | 2 | -11.06
FRUT-1 | H | 3 | -11.23
FRUT-1 | H | 4 | -11.2
FRUT-1 | J | 1 | -11.16
FRUT-1 | J | 2 | -11.24
FRUT-1 | J | 3 | -11.24
FRUT-1 | J | 4 | -11.2
GALT-1 | E | 1 | -21.03
GALT-1 | E | 2 | -21.3
GALT-1 | E | 3 | -21.29
GALT-1 | E | 4 | -21.39
GALT-1 | G | 1 | -21.28
GALT-1 | G | 2 | -21.3
GALT-1 | G | 3 | -21.21
GALT-1 | G | 4 | -21.33
GALT-1 | L | 1 | -21.53
GALT-1 | L | 2 | -21.6
GALT-1 | L | 3 | -21.48
GALT-1 | L | 4 | -21.71

## Lab D
Sample Identity | Sample Name | Replicate Number | Reported d(13C), ‰
IAEA-CH-6 | A | 1 | -10.46
IAEA-CH-6 | A | 2 | -10.43
IAEA-CH-6 | A | 3 | -10.45
IAEA-CH-6 | A | 4 | -10.46
USGS40 | B | 1 | -26.44
USGS40 | B | 2 | -26.44
USGS40 | B | 3 | -26.41
USGS40 | B | 4 | -26.43
USGS62 | C | 1 | -14.8
USGS62 | C | 2 | -14.8
USGS62 | C | 3 | -14.78
USGS62 | C | 4 | -14.81
BEET-1 | D | 1 | -26.14
BEET-1 | D | 2 | -26.04
BEET-1 | D | 3 | -26.13
BEET-1 | D | 4 | -26.06
BEET-1 | I | 1 | -25.98
BEET-1 | I | 2 | -25.95
BEET-1 | I | 3 | -26.01
BEET-1 | I | 4 | -26.01
BEET-1 | K | 1 | -26.06
BEET-1 | K | 2 | -26.1
BEET-1 | K | 3 | -26.06
BEET-1 | K | 4 | -26.11
FRUT-1 | F | 1 | -11
FRUT-1 | F | 2 | -11
FRUT-1 | F | 3 | -10.95
FRUT-1 | F | 4 | -11.2
FRUT-1 | H | 1 | -10.97
FRUT-1 | H | 2 | -10.93
FRUT-1 | H | 3 | -10.95
FRUT-1 | H | 4 | -10.91
FRUT-1 | J | 1 | -10.93
FRUT-1 | J | 2 | -10.94
FRUT-1 | J | 3 | -10.94
FRUT-1 | J | 4 | -10.96
GALT-1 | E | 1 | -21.46
GALT-1 | E | 2 | -21.44
GALT-1 | E | 3 | -21.48
GALT-1 | E | 4 | -21.52
GALT-1 | G | 1 | -21.49
GALT-1 | G | 2 | -21.49
GALT-1 | G | 3 | -21.46
GALT-1 | G | 4 | -21.47
GALT-1 | L | 1 | -21.47
GALT-1 | L | 2 | -21.45
GALT-1 | L | 3 | -21.46
GALT-1 | L | 4 | -21.49

## Lab E
Sample Identity | Sample Name | Replicate Number | Reported d(13C), ‰
IAEA-CH-6 | A | 1 | -10.54
IAEA-CH-6 | A | 2 | -10.38
IAEA-CH-6 | A | 3 | -10.45
IAEA-CH-6 | A | 4 | -10.45
USGS40 | B | 1 | -26.33
USGS40 | B | 2 | -26.43
USGS40 | B | 3 | -26.38
USGS40 | B | 4 | -26.47
USGS62 | C | 1 | -14.81
USGS62 | C | 2 | -14.85
USGS62 | C | 3 | -14.89
USGS62 | C | 4 | -14.83
USGS62 | C | 5 | -14.85
BEET-1 | D | 1 | -26.03
BEET-1 | D | 2 | -25.96
BEET-1 | D | 3 | -26.03
BEET-1 | D | 4 | -26.03
BEET-1 | I | 1 | -26.05
BEET-1 | I | 2 | -26.08
BEET-1 | I | 3 | -25.98
BEET-1 | I | 4 | -26.03
BEET-1 | K | 1 | -26.05
BEET-1 | K | 2 | -26.05
BEET-1 | K | 3 | -26.09
BEET-1 | K | 4 | -26.06
FRUT-1 | F | 1 | -10.97
FRUT-1 | F | 2 | -11.02
FRUT-1 | F | 3 | -11.07
FRUT-1 | H | 1 | -10.93
FRUT-1 | H | 2 | -10.99
FRUT-1 | H | 3 | -10.97
FRUT-1 | H | 4 | -10.98
FRUT-1 | J | 1 | -11.06
FRUT-1 | J | 2 | -10.98
FRUT-1 | J | 3 | -11.02
FRUT-1 | J | 4 | -11.02
GALT-1 | E | 1 | -21.45
GALT-1 | E | 2 | -21.38
GALT-1 | E | 3 | -21.48
GALT-1 | E | 4 | -21.42
GALT-1 | G | 1 | -21.41
GALT-1 | G | 2 | -21.48
GALT-1 | G | 3 | -21.4
GALT-1 | G | 4 | -21.45
GALT-1 | L | 1 | -21.47
GALT-1 | L | 2 | -21.44
GALT-1 | L | 3 | -21.44
GALT-1 | L | 4 | -21.46
GALT-1 | L | 5 | -21.48

## Lab F
Sample Identity | Sample Name | Replicate Number | Reported d(13C), ‰
IAEA-CH-6 | A | 1 | -10.54
IAEA-CH-6 | A | 2 | -10.39
IAEA-CH-6 | A | 3 | -10.42
IAEA-CH-6 | A | 4 | -10.51
USGS40 | B | 1 | -26.33
USGS40 | B | 2 | -26.24
USGS40 | B | 3 | -26.43
USGS40 | B | 4 | -26.39
USGS62 | C | 1 | -14.52
USGS62 | C | 2 | -14.52
USGS62 | C | 3 | -14.79
USGS62 | C | 4 | -14.87
BEET-1 | D | 1 | -25.93
BEET-1 | D | 2 | -26.08
BEET-1 | D | 3 | -25.89
BEET-1 | D | 4 | -26.01
BEET-1 | I | 1 | -25.9
BEET-1 | I | 2 | -25.91
BEET-1 | I | 3 | -26
BEET-1 | I | 4 | -25.99
BEET-1 | K | 1 | -25.98
BEET-1 | K | 2 | -26.13
BEET-1 | K | 3 | -25.99
BEET-1 | K | 4 | -26.03
FRUT-1 | F | 1 | -11.1
FRUT-1 | F | 2 | -11.06
FRUT-1 | F | 3 | -10.64
FRUT-1 | F | 4 | -10.83
FRUT-1 | H | 1 | -11.38
FRUT-1 | H | 2 | -11.2
FRUT-1 | H | 3 | -10.98
FRUT-1 | H | 4 | -11.08
FRUT-1 | J | 1 | -11.04
FRUT-1 | J | 2 | -11.05
FRUT-1 | J | 3 | -10.99
FRUT-1 | J | 4 | -11.03
GALT-1 | E | 1 | -21.49
GALT-1 | E | 2 | -21.43
GALT-1 | E | 3 | -21.44
GALT-1 | E | 4 | -21.47
GALT-1 | G | 1 | -21.25
GALT-1 | G | 2 | -21.26
GALT-1 | G | 3 | -21.29
GALT-1 | G | 4 | -21.37
GALT-1 | L | 1 | -21.27
GALT-1 | L | 2 | -21.51
GALT-1 | L | 3 | -21.43
GALT-1 | L | 4 | -21.55

## Lab G NRC-1
Sample Identity | Sample Name | Replicate Number | Reported d(13C), ‰
IAEA-CH-6 | A | 1 | -10.5
IAEA-CH-6 | A | 2 | -10.43
IAEA-CH-6 | A | 3 | -10.43
IAEA-CH-6 | A | 4 | -10.31
USGS40 | B | 1 | -26.37
USGS40 | B | 2 | -26.43
USGS40 | B | 3 | -26.46
USGS40 | B | 4 | -26.29
USGS62 | C | 1 | -14.89
USGS62 | C | 2 | -14.82
USGS62 | C | 3 | -14.93
USGS62 | C | 4 | -14.72
BEET-1 | D | 1 | -26.03
BEET-1 | D | 2 | -25.97
BEET-1 | D | 3 | -26.04
BEET-1 | D | 4 | -25.93
BEET-1 | I | 1 | -26.07
BEET-1 | I | 2 | -26.17
BEET-1 | I | 3 | -26.17
BEET-1 | I | 4 | -25.97
BEET-1 | K | 1 | -26.06
BEET-1 | K | 2 | -26.08
BEET-1 | K | 3 | -25.94
BEET-1 | K | 4 | -25.94
FRUT-1 | F | 1 | -11.08
FRUT-1 | F | 2 | -10.98
FRUT-1 | F | 3 | -11.02
FRUT-1 | F | 4 | -11.02
FRUT-1 | H | 1 | -11.04
FRUT-1 | H | 2 | -11
FRUT-1 | H | 3 | -11.01
FRUT-1 | H | 4 | -11.03
FRUT-1 | J | 1 | -10.97
FRUT-1 | J | 2 | -10.95
FRUT-1 | J | 3 | -11.01
FRUT-1 | J | 4 | -11
GALT-1 | E | 1 | -21.43
GALT-1 | E | 2 | -21.39
GALT-1 | E | 3 | -21.47
GALT-1 | E | 4 | -21.43
GALT-1 | G | 1 | -21.4
GALT-1 | G | 2 | -21.42
GALT-1 | G | 3 | -21.37
GALT-1 | G | 4 | -21.43
GALT-1 | L | 1 | -21.36
GALT-1 | L | 2 | -21.34
GALT-1 | L | 3 | -21.33
GALT-1 | L | 4 | -21.3

## Lab H NRC-2
Sample Identity | Sample Name | Replicate Number | Reported d(13C), ‰
IAEA-CH-6 | A | 1 | -10.47
IAEA-CH-6 | A | 2 | -10.44
IAEA-CH-6 | A | 3 | -10.5
IAEA-CH-6 | A | 4 | -10.37
USGS40 | B | 1 | -26.38
USGS40 | B | 2 | -26.38
USGS40 | B | 3 | -26.4
USGS40 | B | 4 | -26.29
USGS62 | C | 1 | -14.78
USGS62 | C | 2 | -14.81
USGS62 | C | 3 | -14.83
USGS62 | C | 4 | -14.78
BEET-1 | D | 1 | -26.1
BEET-1 | D | 2 | -26.12
BEET-1 | D | 3 | -26.01
BEET-1 | D | 4 | -26.02
BEET-1 | I | 1 | -26.13
BEET-1 | I | 2 | -26.02
BEET-1 | I | 3 | -26.06
BEET-1 | I | 4 | -26.05
BEET-1 | K | 1 | -25.97
BEET-1 | K | 2 | -26
BEET-1 | K | 3 | -26.08
BEET-1 | K | 4 | -25.94
FRUT-1 | F | 1 | -11
FRUT-1 | F | 2 | -10.97
FRUT-1 | F | 3 | -10.95
FRUT-1 | F | 4 | -11
FRUT-1 | H | 1 | -11.02
FRUT-1 | H | 2 | -11.08
FRUT-1 | H | 3 | -11.11
FRUT-1 | H | 4 | -11.05
FRUT-1 | J | 1 | -10.95
FRUT-1 | J | 2 | -10.96
FRUT-1 | J | 3 | -10.93
FRUT-1 | J | 4 | -10.95
GALT-1 | E | 1 | -21.4
GALT-1 | E | 2 | -21.4
GALT-1 | E | 3 | -21.44
GALT-1 | E | 4 | -21.34
GALT-1 | G | 1 | -21.48
GALT-1 | G | 2 | -21.4
GALT-1 | G | 3 | -21.41
GALT-1 | G | 4 | -21.42
GALT-1 | L | 1 | -21.36
GALT-1 | L | 2 | -21.35
GALT-1 | L | 3 | -21.32
GALT-1 | L | 4 | -21.35

## Lab I NRC-3
Sample Identity | Sample Name | Replicate Number | Reported d(13C), ‰
IAEA-CH-6 | A | 1 | -10.43
IAEA-CH-6 | A | 2 | -10.44
IAEA-CH-6 | A | 3 | -10.5
IAEA-CH-6 | A | 4 | -10.39
USGS40 | B | 1 | -26.4
USGS40 | B | 2 | -26.37
USGS40 | B | 3 | -26.43
USGS40 | B | 4 | -26.41
USGS62 | C | 1 | -14.84
USGS62 | C | 2 | -14.78
USGS62 | C | 3 | -14.83
USGS62 | C | 4 | -14.8
BEET-1 | D | 1 | -26.04
BEET-1 | D | 2 | -25.91
BEET-1 | D | 3 | -26.14
BEET-1 | D | 4 | -26
BEET-1 | I | 1 | -25.98
BEET-1 | I | 2 | -25.97
BEET-1 | I | 3 | -26
BEET-1 | I | 4 | -26.16
BEET-1 | K | 1 | -25.96
BEET-1 | K | 2 | -26.14
BEET-1 | K | 3 | -26.04
BEET-1 | K | 4 | -26.08
FRUT-1 | F | 1 | -10.94
FRUT-1 | F | 2 | -10.9
FRUT-1 | F | 3 | -10.91
FRUT-1 | F | 4 | -10.94
FRUT-1 | H | 1 | -10.98
FRUT-1 | H | 2 | -10.98
FRUT-1 | H | 3 | -10.93
FRUT-1 | H | 4 | -10.93
FRUT-1 | J | 1 | -11
FRUT-1 | J | 2 | -10.97
FRUT-1 | J | 3 | -11
FRUT-1 | J | 4 | -10.95
GALT-1 | E | 1 | -21.37
GALT-1 | E | 2 | -21.4
GALT-1 | E | 3 | -21.4
GALT-1 | E | 4 | -21.39
GALT-1 | G | 1 | -21.4
GALT-1 | G | 2 | -21.44
GALT-1 | G | 3 | -21.42
GALT-1 | G | 4 | -21.42
GALT-1 | L | 1 | -21.46
GALT-1 | L | 2 | -21.44
GALT-1 | L | 3 | -21.42
GALT-1 | L | 4 | -21.42

## Homogeneity
BEET-1 ~ 90 µg (single crystal) 5 replicates from ampule 1720 |  | BEET-1 ~ 180 µg (double crystal) 5 replicates from ampule 1720 |  | BEET-1 550-600 µg |  | GALT-1 ~ 90 µg 5 replicates from ampule 931 |  | GALT-1 ~180 µg 5 replicates from ampule 931 |  | GALT-1 550-600 µg |  | FRUT-1 ~90 µg (single crystal) 5 replicates from ampule 607 |  | FRUT-1 ~180 µg (double crystal) 5 replicates from ampule 607 |  | FRUT-1 550-600 µg
mass (µg) | d(13C), ‰ | mass (µg) | d(13C), ‰ | Ampule Number | d(13C), ‰ | mass (µg) | d(13C), ‰ | mass (µg) | d(13C), ‰ | Ampule Number | d(13C), ‰ | mass (µg) | d(13C), ‰ | mass (µg) | d(13C), ‰ | Ampule Number | d(13C), ‰
94 | -26.13 | 175 | -26.05 | 1 | -26.01 | 86 | -21.11 | 181 | -21.34 | 40 | -21.39 | 88 | -9.98 | 172 | -9.42 | 5 | -10.99
89 | -25.87 | 180 | -26.01 | 1 | -26.04 | 92 | -21.42 | 176 | -21.34 | 40 | -21.35 | 90 | -10.79 | 176 | -11.02 | 5 | -11.04
95 | -26.1 | 177 | -25.81 | 110 | -26.16 | 93 | -21.41 | 179 | -21.38 | 160 | -21.37 | 95 | -10.92 | 181 | -10.95 | 40 | -11
95 | -26.05 | 175 | -26.11 | 110 | -25.86 | 93 | -21.3 | 175 | -21.31 | 160 | -21.38 | 89 | -10.8 | 182 | -10.84 | 40 | -11
94 | -25.96 | 176 | -26.01 | 230 | -25.95 | 92 | -21.28 | 180 | -21.42 | 270 | -21.36 | 95 | -11.03 | 184 | -10.96 | 90 | -10.96
|  |  |  | 230 | -26.01 |  |  |  |  | 270 | -21.35 |  |  |  |  | 90 | -11.01
|  |  |  | 350 | -25.99 |  |  |  |  | 390 | -21.4 |  |  |  |  | 150 | -10.96
|  |  |  | 350 | -25.98 |  |  |  |  | 390 | -21.32 | outlier - DISCARDED |  | outlier - DISCARDED |  | 150 | -10.98
|  |  |  | 480 | -26.06 |  |  |  |  | 500 | -21.37 |  |  |  |  | 202 | -10.98
|  |  |  | 480 | -25.9 |  |  |  |  | 500 | -21.32 |  |  |  |  | 202 | -11
|  |  |  | 610 | -26.01 |  |  |  |  | 620 | -21.35 |  |  |  |  | 260 | -11
|  |  |  | 610 | -25.94 |  |  |  |  | 620 | -21.33 |  |  |  |  | 260 | -11
|  |  |  | 740 | -25.87 |  |  |  |  | 730 | -21.36 |  |  |  |  | 310 | -10.99
|  |  |  | 740 | -26.03 |  |  |  |  | 730 | -21.39 |  |  |  |  | 310 | -11
|  |  |  | 870 | -26.01 |  |  |  |  | 850 | -21.35 |  |  |  |  | 364 | -10.99
|  |  |  | 870 | -26 |  |  |  |  | 850 | -21.36 |  |  |  |  | 364 | -11.03
|  |  |  | 990 | -25.99 |  |  |  |  | 931 | -21.33 |  |  |  |  | 445 | -10.97
|  |  |  | 990 | -26 |  |  |  |  | 931 | -21.31 |  |  |  |  | 445 | -11.02
|  |  |  | 1120 | -26.01 |  |  |  |  | 1013 | -21.35 |  |  |  |  | 500 | -10.91
|  |  |  | 1120 | -25.98 |  |  |  |  | 1013 | -21.4 |  |  |  |  | 500 | -10.95
|  |  |  | 1250 | -25.97 |  |  |  |  | 1093 | -21.36 |  |  |  |  | 550 | -10.94
|  |  |  | 1250 | -26.03 |  |  |  |  | 1093 | -21.38 |  |  |  |  | 550 | -11.03
|  |  |  | 1370 | -26.02 |  |  |  |  | 1175 | -21.27 |  |  |  |  | 607 | -10.97
|  |  |  | 1370 | -25.94 |  |  |  |  | 1175 | -21.34 |  |  |  |  | 607 | -10.99
|  |  |  | 1480 | -26 |  |  |  |  | 1255 | -21.33 |  |  |  |  | 660 | -10.98
|  |  |  | 1480 | -26.03 |  |  |  |  | 1255 | -21.33 |  |  |  |  | 660 | -10.99
|  |  |  | 1560 | -26.01 |  |  |  |  | 1336 | -21.33 |  |  |  |  | 710 | -10.98
|  |  |  | 1560 | -26.11 |  |  |  |  | 1336 | -21.37 |  |  |  |  | 710 | -11.04
|  |  |  | 1640 | -26.012 |  |  |  |  | 1418 | -21.31 |  |  |  |  | 769 | -10.96
|  |  |  | 1640 | -26.05 |  |  |  |  | 1418 | -21.31 |  |  |  |  | 769 | -11.02
|  |  |  | 1720 | -26.05 |  |  |  |  | 1498 | -21.3 |  |  |  |  | 850 | -10.97
|  |  |  | 1720 | -26.13 |  |  |  |  | 1498 | -21.41 |  |  |  |  | 850 | -10.99
|  |  |  | 1800 | -26.01 |  |  |  |  | 1579 | -21.33 |  |  |  |  | 910 | -10.93
|  |  |  | 1800 | -25.95 |  |  |  |  | 1579 | -21.38 |  |  |  |  | 910 | -11
|  |  |  | 1880 | -25.98 |  |  |  |  | 1661 | -21.33 |  |  |  |  | 960 | -11
|  |  |  | 1880 | -26.09 |  |  |  |  | 1661 | -21.39 |  |  |  |  | 960 | -10.99
|  |  |  | 1960 | -25.95 |  |  |  |  | 1741 | -21.32 |  |  |  |  | 1012 | -10.96
|  |  |  | 1960 | -26.11 |  |  |  |  | 1741 | -21.41 |  |  |  |  | 1012 | -11.04
|  |  |  | 1995 | -25.93 |  |  |  |  | 1822 | -21.35 |  |  |  |  | 1093 | -10.97
|  |  |  | 1995 | -26.09 |  |  |  |  | 1822 | -21.39 |  |  |  |  | 1093 | -11.03

## Thermal Stability
BEET-1 |  |  | GALT-1 |  |  | FRUT-1
Temperature, ° C | Ampule | d(13C), ‰ | Temperature, ° C | Ampule | d(13C), ‰ | Temperature, ° C | Ampule | d(13C), ‰
37 | 944 | -25.96 | 37 | 874 | -21.44 | 37 | 523 | -10.95
37 | 944 | -26 | 37 | 874 | -21.46 | 37 | 523 | -10.94
37 | 944 | -25.98 | 37 | 874 | -21.46 | 37 | 523 | -10.97
37 | 945 | -26.01 | 37 | 875 | -21.45 | 37 | 524 | -10.97
37 | 945 | -26.02 | 37 | 875 | -21.47 | 37 | 524 | -10.99
37 | 945 | -25.95 | 37 | 875 | -21.43 | 37 | 524 | -10.95
21.8 | 946 | -25.93 | 21.5 | 876 | -21.43 | 21.7 | 525 | -10.98
21.8 | 946 | -25.98 | 21.5 | 876 | -21.44 | 21.7 | 525 | -10.93
21.8 | 946 | -26.08 | 21.5 | 876 | -21.39 | 21.7 | 525 | -10.94
21.8 | 947 | -25.99 | 21.5 | 877 | -21.41 | 21.7 | 526 | -10.96
21.8 | 947 | -25.93 | 21.5 | 877 | -21.43 | 21.7 | 526 | -10.95
21.8 | 947 | -25.99 | 21.5 | 877 | -21.4 | 21.7 | 526 | -10.99
-28 | 948 | -26.02 | -28 | 878 | -21.45 | -28 | 527 | -10.96
-28 | 948 | -25.95 | -28 | 878 | -21.42 | -28 | 527 | -10.92
-28 | 948 | -26.03 | -28 | 878 | -21.42 | -28 | 527 | -10.95
-28 | 949 | -25.97 | -28 | 879 | -21.45 | -28 | 528 | -11
-28 | 949 | -25.98 | -28 | 879 | -21.44 | -28 | 528 | -10.97
-28 | 949 | -26.02 | -28 | 879 | -21.44 | -28 | 528 | -11

## Temporal Stability
BEET-1 Ampule 1720 |  | GALT-1 Ampule 931 |  | FRUT-1 Ampule 607
Date | d(13C), ‰ | Date | d(13C), ‰ | Date | d(13C), ‰
October 2016 | -26.03 | October 2016 | -21.4 | October 2016 | -11.04
October 2016 | -25.97 | October 2016 | -21.44 | October 2016 | -11
October 2016 | -26.04 | October 2016 | -21.42 | October 2016 | -11.01
October 2016 | -25.93 | October 2016 | -21.42 | October 2016 | -11.03
January 2018 | -25.97 | January 2018 | -21.48 | January 2018 | -11.03
January 2018 | -26.05 | January 2018 | -21.42 | January 2018 | -11.04
January 2018 | -26.06 | January 2018 | -21.45 | January 2018 | -10.87
February 2018 (1) | -26 | February 2018 (1) | -21.42 | February 2018 (1) | -10.97
February 2018 (1) | -26.05 | February 2018 (1) | -21.41 | February 2018 (1) | -11.03
February 2018 (1) | -26 | February 2018 (1) | -21.44 | February 2018 (1) | -11
February 2018 (2) | -26.01 | February 2018 (2) | -21.44 | February 2018 (2) | -11.02
February 2018 (2) | -26 | February 2018 (2) | -21.4 | February 2018 (2) | -11.01
February 2018 (2) | -25.94 | February 2018 (2) | -21.47 | February 2018 (2) | -10.98

## Consistency check of other CRMs
Data used to assess the coherence of the assigned carbon isotope delta values in the set of eleven international CRMs (Figure 1)
δref(13C) is the isotope delta value of the various materials expressed against the in-house internal CO2 reference gas (with the assigned value of 0 for both carbon and oxygen)
Run | Material | Replicate | δref(13C), ‰
1 | IAEA-600 | 1 | -8.2148
1 | IAEA-600 | 2 | -8.1739
1 | IAEA-600 | 3 | -8.2092
1 | IAEA-CH-3 | 1 | -5.2259
1 | IAEA-CH-3 | 2 | -5.2435
1 | IAEA-CH-3 | 3 | -5.2002
1 | IAEA-CH-6 | 1 | 9.1727
1 | IAEA-CH-6 | 2 | 9.2048
1 | IAEA-CH-6 | 3 | 9.2078
1 | IAEA-CH-7 | 1 | -12.6639
1 | IAEA-CH-7 | 2 | -12.7239
1 | IAEA-CH-7 | 3 | -12.7005
1 | USGS40 | 1 | -6.7972
1 | USGS40 | 2 | -6.8058
1 | USGS40 | 3 | -6.7419
1 | USGS61 | 1 | -15.6056
1 | USGS61 | 2 | -15.5085
1 | USGS61 | 3 | -15.5874
1 | USGS62 | 1 | 4.8171
1 | USGS62 | 2 | 4.8694
1 | USGS62 | 3 | 4.873
1 | USGS63 | 1 | 18.4974
1 | USGS63 | 2 | 18.5795
1 | USGS63 | 3 | 18.5399
1 | USGS64 | 1 | -21.4324
1 | USGS64 | 2 | -21.3889
1 | USGS64 | 3 | -21.4166
1 | USGS65 | 1 | -0.71065
1 | USGS65 | 2 | -0.71859
1 | USGS65 | 3 | -0.67209
1 | USGS66 | 1 | 19.0675
1 | USGS66 | 2 | 19.0865
1 | USGS66 | 3 | 19.072
2 | IAEA-600 | 1 | -8.0857
2 | IAEA-600 | 2 | -8.1385
2 | IAEA-600 | 3 | -8.0956
2 | IAEA-CH-3 | 1 | -5.1269
2 | IAEA-CH-3 | 2 | -5.1119
2 | IAEA-CH-3 | 3 | -5.1075
2 | IAEA-CH-6 | 1 | 9.3304
2 | IAEA-CH-6 | 2 | 9.363
2 | IAEA-CH-6 | 3 | 9.3478
2 | IAEA-CH-7 | 1 | -12.5664
2 | IAEA-CH-7 | 2 | -12.6458
2 | IAEA-CH-7 | 3 | -12.584
2 | USGS40 | 1 | -6.7282
2 | USGS40 | 2 | -6.708
2 | USGS40 | 3 | -6.7114
2 | USGS61 | 1 | -15.4794
2 | USGS61 | 2 | -15.4514
2 | USGS61 | 3 | -15.4673
2 | USGS62 | 1 | 5.0033
2 | USGS62 | 2 | 4.9508
2 | USGS62 | 3 | 4.9376
2 | USGS63 | 1 | 18.6848
2 | USGS63 | 2 | 18.6371
2 | USGS63 | 3 | 18.6612
2 | USGS64 | 1 | -21.2879
2 | USGS64 | 2 | -21.355
2 | USGS64 | 3 | -21.3362
2 | USGS65 | 1 | -0.68916
2 | USGS65 | 2 | -0.6113
2 | USGS65 | 3 | -0.63652
2 | USGS66 | 1 | 19.206
2 | USGS66 | 2 | 19.1817
2 | USGS66 | 3 | 19.1798
3 | IAEA-600 | 1 | -8.1159
3 | IAEA-600 | 2 | -8.1579
3 | IAEA-600 | 3 | -8.1627
3 | IAEA-CH-3 | 1 | -5.1716
3 | IAEA-CH-3 | 2 | -5.1462
3 | IAEA-CH-3 | 3 | -5.119
3 | IAEA-CH-6 | 1 | 9.2516
3 | IAEA-CH-6 | 2 | 9.2978
3 | IAEA-CH-6 | 3 | 9.2967
3 | IAEA-CH-7 | 1 | -12.654
3 | IAEA-CH-7 | 2 | -12.6592
3 | IAEA-CH-7 | 3 | -12.6548
3 | USGS40 | 1 | -6.752
3 | USGS40 | 2 | -6.7785
3 | USGS40 | 3 | -6.8024
3 | USGS61 | 1 | -15.5348
3 | USGS61 | 2 | -15.5545
3 | USGS61 | 3 | -15.5542
3 | USGS62 | 1 | 4.9242
3 | USGS62 | 2 | 4.8944
3 | USGS62 | 3 | 4.9004
3 | USGS63 | 1 | 18.6654
3 | USGS63 | 2 | 18.6301
3 | USGS63 | 3 | 18.6019
3 | USGS64 | 1 | -21.3422
3 | USGS64 | 2 | -21.4019
3 | USGS64 | 3 | -21.3761
3 | USGS65 | 1 | -0.67745
3 | USGS65 | 2 | -0.66075
3 | USGS65 | 3 | -0.66978
3 | USGS66 | 1 | 19.1357
3 | USGS66 | 2 | 19.1451
3 | USGS66 | 3 | 19.1328

## 17O correction
The effect of the choice of oxygen-17 correction model (Lab G NRC-1 data)
δ(13C) values normalized against four international CRMs (IAEA-CH-6, IAEA-CH-7, USGS40, USGS62)
|  | 17O correction method |  | Difference between 17O correction methods
Sample | ID | SSH | IUPAC | SSH-IUPAC
IAEA-CH-6 | A-1 | -10.5019 | -10.5012 | -0.0006813
IAEA-CH-6 | A-2 | -10.4275 | -10.4276 | 0.00016555
IAEA-CH-6 | A-3 | -10.4333 | -10.4336 | 0.00030922
IAEA-CH-6 | A-4 | -10.3143 | -10.3146 | 0.00022825
USGS40 | B-1 | -26.3706 | -26.3706 | 3.1652e-05
USGS40 | B-2 | -26.4252 | -26.426 | 0.00080081
USGS40 | B-3 | -26.4627 | -26.4634 | 0.00066096
USGS40 | B-4 | -26.2876 | -26.2883 | 0.00074103
USGS62 | C-1 | -14.8904 | -14.8904 | 4.1475e-05
USGS62 | C-2 | -14.8212 | -14.8219 | 0.00065815
USGS62 | C-3 | -14.9334 | -14.934 | 0.0005637
USGS62 | C-4 | -14.724 | -14.7246 | 0.00052612
BEET-1 | D-1 | -26.0257 | -26.0255 | -0.00020691
BEET-1 | D-2 | -25.9735 | -25.9734 | -3.9082e-05
BEET-1 | D-3 | -26.0436 | -26.0437 | 8.6845e-05
BEET-1 | D-4 | -25.9316 | -25.9316 | 5.9154e-05
GALT-1 | E-1 | -21.4318 | -21.4319 | 9.1758e-05
GALT-1 | E-2 | -21.387 | -21.387 | -1.8654e-05
GALT-1 | E-3 | -21.4671 | -21.4671 | 9.4193e-05
GALT-1 | E-4 | -21.4255 | -21.4255 | 2.7171e-05
FRUT-1 | F-1 | -11.077 | -11.0773 | 0.0002588
FRUT-1 | F-2 | -10.9763 | -10.9766 | 0.0003368
FRUT-1 | F-3 | -11.0192 | -11.0197 | 0.00042762
FRUT-1 | F-4 | -11.018 | -11.0183 | 0.00032634
GALT-1 | G-1 | -21.3981 | -21.3985 | 0.00035244
GALT-1 | G-2 | -21.4211 | -21.4215 | 0.00039965
GALT-1 | G-3 | -21.3742 | -21.3745 | 0.0002919
GALT-1 | G-4 | -21.4287 | -21.4291 | 0.00038214
FRUT-1 | H-1 | -11.0408 | -11.0411 | 0.00029755
FRUT-1 | H-2 | -10.9973 | -10.9977 | 0.0003761
FRUT-1 | H-3 | -11.0119 | -11.0123 | 0.00039023
FRUT-1 | H-4 | -11.0253 | -11.0255 | 0.00024827
BEET-1 | I-1 | -26.0685 | -26.069 | 0.00044713
BEET-1 | I-2 | -26.1654 | -26.1659 | 0.00052981
BEET-1 | I-3 | -26.1697 | -26.17 | 0.00035671
BEET-1 | I-4 | -25.9728 | -25.9732 | 0.00039711
FRUT-1 | J-1 | -10.9729 | -10.9733 | 0.00037743
FRUT-1 | J-2 | -10.9546 | -10.955 | 0.00033592
FRUT-1 | J-3 | -11.0082 | -11.0085 | 0.00029218
FRUT-1 | J-4 | -10.9957 | -10.996 | 0.00026398
BEET-1 | K-1 | -26.0617 | -26.062 | 0.00032561
BEET-1 | K-2 | -26.0771 | -26.0775 | 0.0004197
BEET-1 | K-3 | -25.9353 | -25.9357 | 0.00039987
BEET-1 | K-4 | -25.9352 | -25.9356 | 0.0003384
GALT-1 | L-1 | -21.3574 | -21.3577 | 0.00032181
GALT-1 | L-2 | -21.3409 | -21.3412 | 0.00035482
GALT-1 | L-3 | -21.3281 | -21.3285 | 0.00036414
GALT-1 | L-4 | -21.3029 | -21.3032 | 0.00028663